## Sheet1
| SMP3 | SMP11
| SMP3 measures global solar radiation on a horizontal plane. | SMP11 measures incoming global solar radiation with a 180° field of view.
Spectral range (50% points) | 300 to 2800 nm | 285 to 2800 nm
Response time (63%) | < 1.5 s | < 0.7 s
Response time (95%) | < 12 s | < 2 s
Zero offset A | < 15 W/m² | < 7 W/m²
Zero offset B | < 5 W/m² | < 2 W/m²
Directional response (up to 80° with 1000 W/m² beam) | 20 W/m² | < 10 W/m²
Temperature dependence of sensitivity (-20 ºC to +50 ºC) | < 3 % | < 1 %
Analogue output (-V version) | 0 to 1 V | 0 to 1 V
Analogue output (-A version) | 4 to 20 mA | 4 to 20 mA
Digital output | 2-wire RS-485 | 2-wire RS-485
Classification to ISO 9060:1990 | Second Class | Secondary Standard
Analogue output • V-version | 0 to 1V | 0 to 1V
Analogue output range* | -200 to 2000W/m | -200 to 2000W/m
Analogue output • A-version | 4 to 20mA | 4 to 20mA
Analogue output range* | 0 to 1600W/m | 0 to 1600W/m
Serial output | RS-485 Modbus® | RS-485 Modbus®
Serial output range | -400 to 2000W/m | -400 to 4000W/m
Response time (63%) | < 1.5 s | < 0.7 s
Response time (95%) | < 12 s | < 2 s
Spectral range (20 % points) | 285 to 3000nm | 270 to 3000nm
Spectral range (50 % points) | 300 to 2800nm | 285 to 2800nm
Zero o sets (unventilated) | < 15W/m | < 7 W/m
(a) thermal radiation (at 200 W/m ) | < 5W/m | < 2 W/m
(b) temperature change (5 K/h) | < 1% | < 0.5%
Non-stability (change/year) | < 1.5% | < 0.2%
Non-linearity (100 to 1000W/m ) | < 20W/m | < 10 W/m
Directional response | < 2% (-20 °C to +50 °C) | < 1% (-20°C to +50 °C)
(up to 80 ° with 1000W/m beam) | < 4% (-40 °C to +70 °C) | < 2% (-40 °C to +70 °C)
Temperature response | < 1% | < 1%
Spectral selectivity (350 to 1500nm) | < 1% | < 0.2%
Tilt response (0 ° to 90 ° at 1000W/m ) | 180 ° | 180 °
Field of view | < 0.2 ° | < 0.1 °
Accuracy of bubble level | V-version: 55mW | V-version: 55mW
Power consumption (at 12 VDC) | A-version: 100mW | A-version: 100mW
Supply voltage | 5 to 30VDC | 5 to 30VDC
Detector type | Thermopile | Thermopile
Operating and storage temperature range | -40 °C to +80 °C | -40 °C to +80 °C
Humidity range | 0 to 100% | 0 to 100%
MTBF (Mean Time Between Failures) ** | > 10 years | > 10 years
Ingress Protection (IP) rating | 67 | 67
So ware, Windows™